WATER MONITORING

# QUALITY ASSURANCE PROJECT PLAN

# Title and Approval Page

Yuba River Monitoring Program
QUALITY ASSURANCE PROJECT PLAN

refer correspondence to:

Dominic Gregorio

Division of Water Quality

State Water Resources Control Board

901 P St.

Sacramento CA 95814

OUTLINE PREPARED BY:

First Edition: Gwen Starrett, SWRCB, 1998

Second Edition: Dominic Gregorio, SWRCB, 2000

COMPLETED PLAN PREPARED BY:

Maureen Rose, South Yuba River Citizens League, August 2000

Approvals:

United States Environmental Protection Agency, Region IX

Signature:  ________________________________________________Date: _______________________

State Water Resources Control Board

Signature:  ________________________________________________Date: _______________________

State Water Resources Control Board

Signature:  ________________________________________________Date: _______________________

Regional Water Quality Control Board

Signature:  ________________________________________________Date: _______________________

# 2.	Table of Contents

1. Title and Approval Page

2. Table of Contents

3. Distribution List

4. Project/Task Organization

5. Problem Definition/Background

5.1. Problem Statement

5.2. Intended Usage of Data

6. Project/Task Description

6.1. General Overview of Project

6.2. Project Timetable

7. Measurement Quality Objectives

7.1. Accuracy

7.2. Standardization of Instruments and Test Procedures

7.3. Comparability

7.4. Completeness

7.5. Precision

7.6. Representativeness

7.7. Sensitivity

8. Training Requirements and Certification

9. Documentation and Records

10. Sampling Process Design

10.1. Rationale for Selection of Sampling Sites

10.2. Sample Design Logistics

11. Sampling Method Requirements

12. Sample Handling and Custody Procedures

12.1. Sample Handling

12.2. Custody Procedures

12.3. Disposal

13. Analytical Methods Requirements

14. Quality Control Requirements

14.1. Cautions regarding  Test Procedures

14.2. Field Blanks, Replicate Samples and Split Samples

15. Instrument/Equipment Testing, Inspection and Maintenance Requirements

15.1. Temperature

15.2. Dissolved Oxygen

15.3. pH and Conductivity

15.4. Turbidity

15.5. Nutrients and Urban Pollutants

16. Instrument Calibration and Frequency

17. Inspection/Acceptance Requirements

18. Data Acquisition Requirements

18.1. Analytical Data

18.2. Geographical Information/ Mapping

19. Data Management

20. Assessment and Response Actions

21. Reports

22. Data Review, Validation and Verification

23. Validation and Verification Methods

24. Reconciliation with DQOs

# 3.	Distribution List

All group leaders, and technical advisors will receive copies of this Quality Assurance (QA) plan, and any approved revisions of this plan.  Once approved, this QA plan will be available to any interested party by requesting a copy from Dominic Gregorio (see address on title page).

# 4.	Project Organization

The Yuba River Monitoring Program is a citizezen based monitoring program where local residents will

monitor and assess the Yuba River watershed.

Several resource agencies have assisted in the development of  this project from its conception.  Additional partnerships will be developed to ensure adequate technical support to all participating citizen monitoring groups.  The QA plan reflects the diversity of monitoring and organizational support involved in this project. For the elements of this QA plan, we have addressed aspects that are shared with all groups as well as those aspects that are unique to individual groups.  While the goals of monitoring may vary, the data quality objectives are consistent allowing  us to compare data collected by different organizations.

## Technical Advisory Committee

Dominc Gregorio - Regional Coordinator, State Water Resources Control Board

Marilyn Murphy -   California Department of Parks and Recreation

Lora Konde  - California Department of Fish and Game

Fraser Shilling – UC Davis, Sierra Nevada Network

Chad McDonald , Bear Creek Keeper

John VanderVeen, Friends of Deer Creek

Joanne Hild, Deer Creek Keeper

# 5.	Problem Definition/Background

## 5.1.	Problem Statement

There is insufficient information to adequately assess the status of aquatic resources in the Yuba watershed.  Citizen monitoring organizations have been formed in local watersheds to address their own water quality concerns.  If quality assurance is adequate, valuable information will be provided for watershed management and pollution prevention.

### 5.1.1.	Regional Citizen Monitoring Mission and Goals

#### 5.1.1.1.	Mission

The mission of citizen monitoring is to produce environmental information which is needed to protect California’s watersheds and aquatic resources. Citizen monitoring will also inform and engage the community in effective watershed stewardship.

#### 5.1.1.2.	Watershed Goals

The goals of citizen monitoring are:

- Identifying valued resources and watershed characteristics for setting management goals,

- Identifying physical watershed characteristics influencing pollutant inputs, transport and fate,

- Identifying the status and trends of biological resources in and around an aquatic environment,

- Screening for water quality problems,

- Identifying pollution sources and illegal activities (spills, wetland fill, diversions, discharges),

- Establishing trends in water quality for waters that would otherwise be un-monitored,

- Evaluating the effectiveness of restoration or management practices,

- Evaluating the effect of a particular activity or structure, and

- Evaluating the quality of water compared to specific water quality criteria.

In addition, citizen monitors build awareness of water quality issues, aquatic resources and pollution prevention.

This project will supplement existing agency information by monitoring streams in the Yuba watershed.  The focus of the project is on habitat and chemical, physical and biological water quality measures that will identify the status of these aquatic resources.  Analyses, for the most part, will be conducted in the field with test kits and field instruments.  This information will be provided to the regulatory agencies.

The following paragraphs identify the goals of the Yuba River Monitoring Program.  The primary goals of the Yuba River Monitoring Project are:

- Design and execute scientifically credible studies which assess the condition of the Yuba River ecosystem.

- Empower citizens to be responsible stewards and decision-makers.

Further goals include:

- Identify valued resources and watershed characteristics for setting management goals,

- Identify physical watershed characteristics influencing pollutant inputs, transport and fate,

- Identify the status and trends of biological resources in and around an aquatic environment,

- Screen for water quality problems,

- Identify pollution sources and potentially illegal activities (spills, wetland fill, diversions, discharges),

- Establish trends in water quality for waters that would otherwise be un-monitored,

- Evaluate the effectiveness of restoration or management practices,

- Evaluate the effect of a particular activity or structure, and

- Evaluate the quality of water compared to specific water quality criteria.

## 5.2.	Intended Usage of Data

Data will be compiled by SYRCL.  The information will be collated and shared with the State Water Resources Control Board, the Central Valley Regional Water Quality Control Board, and upon request to other state, federal, and local agencies and organizations.  The main database will be maintained at a location to be determined by SYRCL, the Deer Creek Monitoring Program, the Bear River Monitoring Program, and the participants in the Yuba Watershed Council.

# 6.	Project/Task Description

The South Yuba River Citizens League is monitoring water quality in the Yuba River watershed.  Physical, chemical and biological parameters will be measured.

This QA plan addresses data quality objectives for the following parameters, which will be sampled through the Yuba River Monitoring Program:

Temperature

Dissolved Oxygen

pH

Conductivity

Suspended Sediment (Turbidity)

Nitrate (nitrogen)

Mercury

Copper

Zinc

Arsenic

Fecal (E. coli) Coliform Bacteria

Oil and Grease

Benthic Macroinvertebrates

For stream and urban storm drain environments flow will be determined by using the protocol described in the U.S. EPA Volunteer Stream Monitoring Manual and/or in the Yuba River Monitoring Program Volunteer Manual.

This program has a systematic method for visual and other sensory observations. A Stream/Shore Walk Visual Assessment observation sheet, with instructions, is included in the Yuba River Monitoring Program Volunteer Manual.  Observations using the Stream/ShoreWalk Visual Assessment observation sheet will be made, at a minimum, on a quarterly basis.  In addition, the stream habitat quality will be assessed, at least once per year, using the California Dept. of Fish and Game Physical Habitat Assessment Form.  Observational data include epifaunal substrate/available cover, embeddedness, velocity/depth regimes, sediment deposition, channel flow status, channel alteration, frequency of riffles, bank stability, vegetative protection, and riparian vegetative zone width.

Analysis for the following parameters are not addressed in this QA plan.  Samples for these and other parameters will be taken by volunteers and staff, and sent to an agency, commercial, or academic laboratory for analysis:

Total Organic Carbon

PAH’s

Pesticides and other synthetic organic compounds

Toxicity

Table 6.1  Type and Frequency of Monitoring in the Yuba River Citizen Monitoring Program

|  |  |  |  | Water Quality Standard Available | Agency or Historical Data Available* |
| --- | --- | --- | --- | --- | --- |
| Temperature | M |  |  | Yes | Yes |
| Dissolved Oxygen | M |  |  | Yes | Yes |
| pH | M |  |  | Yes | Yes |
| Conductivity Salinity | M |  |  |  | Yes |
| Turbidity | M |  |  | Yes | ? |
| Ammonia | M |  |  | Yes | ? |
| Nitrate | M |  |  | Yes | Yes |
| Copper | M |  |  | Yes | ? |
| Metals | M |  |  | Yes | Yes |
| Petroleum | M |  |  | Yes | Yes |
| Bacteria | M |  |  | Yes | Yes |
| Benthic Macroinvertebrates | Bi-Annually |  |  | No | Yes |
| Color | Quarterly |  |  | No |  |
| Trash | X |  |  | No |  |
| Dumping/Spills | X |  |  | No |  |

Frequency:   M: Monthly,   S:  Seasonal, depending on flows, X:  Irregular

*Data not necessarily complete for all parameters.

## General Overview of Project

The Yuba River Monitoring Program was created to provide the monitoring elements required in the coordinated watershed plan for the Yuba River Basin, as funded under Proposition 204. The watershed is composed of the interacting landscapes and river systems. This plan describes procedures for assessing land use/land cover and impacts of particular water quality stressors. By monitoring conditions in both aquatic and terrestrial environments, the “health” of the watershed can be periodically determined relative to standards for water quality and land cover disturbance. “Watershed health” in this case refers to the relative state of the combined landscape and river systems in terms of maintenance of natural ecological, geological, and hydrological processes. The performance standards for a watershed will depend on a combination of legal minima and regional social expectations for ecosystem services and aesthetics.

## 6.2.	Project Timetable

Through outreach meetings and member mailings, SYRCL collected a list of 60 volunteers.  These volunteers signed an agreement to work for the Yuba River Monitoring Project for one full year.  Each volunteer understands that they are responsible for sampling at their assigned site on the assigned monitoring date.  To date, volunteers have been through  five Monitoring trainings:

June 14th, 2000 – Volunteer Monitoring Organizational Meeting.  Volunteers were briefed on the program.  The agenda for the meeting included a review of all sites to be monitored and why the sites were chosen.  Participants learned about the expectations of a Volunteer Monitor and had the opportunity to sign up for a one year term.  Volunteers signed up for specific sites based on their background, experience, and desired geographic region for monitoring.

July 18th and July 24th,  2000 – Volunteer Monitor Trainings – SYRCL staff and selected Technical Advisory Committee members conducted a water quality sampling training.  In this training, volunteers learned about the sampling techniques that they will  use as required under the Quality Assurance Program Plan (see Section 10 - Quality Assurance Policy and Protocols).  Two trainings were held to ensure that all Volunteer Monitors were able to attend the training before the First Sampling Day.

August 5, 2000 - First Sampling Training Day – On the first day of sampling, volunteers met at the South Yuba River near Bridgeport to spend the day learning specific sampling and measurement protocols.  Volunteers received packets with their specific site assignments, sampling protocols, field data sheets and more.  Samples were taken at the site at Bridgeport for training purposes.  Fraser Shilling of UC Davis and Maureen Rose of SYRCL conducted  the training.

August 30, 2000 – An evening refresher course was given on protocols, conducted by Maureen Rose of SYRCL.

September 9, 2000 – Our first day of independent sampling for volunteers.  Anon-river protocol review was given in the morning before sampling.  Data collected is not to be stored with long term data.  However, results will be analyzed for the purpose of conducting quality assurance.

September 26, 2000 – An evening refresher course will be given on protocols, conducted by Maureen Rose of SYRCL.

October 7, 2000 – First Sampling Day – Training on conducting a StreamWalk Survey will occur before sampling.  A refresher course will be given again on protocols.  Actual data collection begins on this date.

Fall/Winter, 2000/2001 – Two quality assurance/quality control trainings will be held for Regional Coordinators/Leaders.

Spring, 2001 – Benthic Macroinvertebrate Training.  An intensive training in the collection procedures for Benthic Macroinvertebrates will be given.

# 7	Measurement Quality Objectives

This section identifies how accurate, precise, complete, comparable, sensitive and representative our measurements will be.  Objectives for these data characteristics are summarized in the Tables 7-1 to 7-5.  Data quality objectives were derived by reviewing the QA plans and performance of other citizen monitoring organizations’ (e.g. Chesapeake Bay, Texas Watch, Coyote Creek Riparian Station, Southern California Citizen Monitoring Steering Committee, Heal the Bay Malibu StreamTeam).

Table 7.1.  Data Quality Objectives for Conventional Water Quality Parameters

| Parameter | Method/range | Units | Detection Limit | Sensitivity | Precision | Accuracy | Completeness |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Temperature | Thermometer (-5 to 50) | o C | -5 | 0.5 o C | ± 10% | ± 10% | 80% |
| Dissolved oxygen | Electronic meter/probe | mg/l | <0.1 | 0.1 mg/l | ± 10% | ± 10% | 80% |
| Dissolved oxygen | Micro-Winkler Titration | mg/l | <0.2 | 0.2 mg/l | ± 10% | ± 10% | 80% |
| pH | pH meter | pH units | 2 | 0.1 unit | ± 10% | ± 10% | 80% |
| Conductivity | conductivity meter | (µS/cm) | 10 | 10 (µS/cm) | ± 10% | ± 10% | 80% |
| Turbidity | Nephelometer | NTU’s | <0.1 | 0.1 | ± 10% | ± 10% | 80% |
| Suspended Sediment | filter and weighing | mg/l | < 0.5 | 0.1 mg/l | ± 10% | ± 10% | 80% |

NA:  not applicable

Table 7.2. Data Quality Objectives for Nutrients Using Colorimeters or Spectrophotometers

| Parameter | Method/range | Units | Detection Limit | Sensitivity* | Precision | Accuracy | Completeness |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Ammonia Nitrogen | Nessler method | mg/l | <0.01 | 0.01 | ± 10% | ± 10% | 80% |
| Nitrate Nitrogen | Cadmium reduction | mg/l | <0.01 | 0.01 | ± 10% | ± 10% | 80% |
| Ortho-Phosphate | Ascorbic acid | mg/l | <0.01 | 0.01 | ± 10% | ± 10% | 80% |

Table 7.3. Data Quality Objectives for Nutrients Using Comparators

| Parameter | Method/range | Units | Detection Limit | Sensitivity* | Precision | Accuracy | Completeness |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Ammonia Nitrogen | Salicylate method | mg/l | <0.25 | 0.25 (0-0.5) 0.5 (0.5-1.0) 1.0 (1.0-2.0) 2.0 (2.0-4.0) | ± 0.5 | NA | 80% |
| Nitrate Nitrogen | Zinc reduction | mg/l | 1.0 | 1.0 | ± 5 | NA | 80% |
| Ortho-Phosphate | Ascorbic acid | mg/l | <0.2 | 0.2 (0-1.0) 0.5 (1.0-2.0) | ± 0.5 | NA | 80% |

Table 7.4. Data Quality Objectives for Urban Pollutants Using Professional Laboratory

| Parameter | Method/range | Units | Detection Limit | Sensitivity* | Precision | Accuracy | Completeness |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Mercury | AA Cold Vapor Technique – EPA Method 245.1 | microg/l | 0.2 | 0.2 | <5% | ±15% | 80% |
| Zinc | EPA Method 289.1 | microg/l | 20 | 20 | <10% | ±10% | 80% |
| Arsenic | EPA Method 200.9 | microg/l | 2 | 2 | <15% | ±15% | 80% |
| Total Copper | Neocuproine (0.25 - 4.0) | microg/l | 0.25 | 0.25 (0-0.5) 0.5 (0.5-2.0) 1.0 (2.0-4.0) | ± 0.5 | NA | 80% |
| Oil and Grease | EPA Method 413.1 | mg/l | 4 | 4 | <10% | ±30% | 80% |

Table 7.5. Data Quality Objectives for Biological Parameters

| Parameter | Method/range | Units | Detection Limit | Sensitivity* | Precision | Accuracy | Completeness |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Benthic Macro-invertebrates | Calif. Stream Bioassessment Protocol | N/A | Family level | N/A | < 5% difference | < 5% difference | 80% |
| Fecal Coliform Bacteria | Colilert 18 hour | MPN/ 100 ml | 10 | See IDEXX quantitray tables | Duplicates within ½ of an order of magnitude | Positive standard within ½ of an order of magnitude | 80% |

## 7.1.	Accuracy

Accuracy describes how close the measurement is to its true value.  Accuracy is the measurement of a sample of known concentration and comparing the known value against the measured value.  The accuracy of chemical measurements will be checked by performing tests on standards at the quality control sessions held twice a year.  A standard is a known concentration of a certain solution.  Standards can be purchased from chemical or scientific supply companies.  Standards might also be prepared by a professional partner, e.g. a local analytical laboratory, certified for water or wastewater analysis by EPA.  The concentration of the standards, known to the volunteer leader, will be unknown to the monitors until after measurements are determined.  The concentration of the standards should be within the mid-range of the equipment.

For all chemical water quality parameters volunteers should obtain results within 10% of the true value, when the true value is within the mid-range of the expected values.

Because of the insensitivity of some of the nutrient (using comparators) and turbidity kits, an accuracy of 10% may not be feasible. For the nutrient tests using comparators, the results of all tests should be within one reading of the true value.  For example, all analyses of an ammonia standard of 0.6 mg/l should be either 0.5 or 1.0 mg/l.

The nitrate test will measure nitrite as well as nitrate.  At the QC sessions, we will sample surface waters and test for nitrate.

Accuracy for bacterial parameters will be determined by analyzing a positive control sample.  A positive control is similar to a standard, except that a specific discreet value is not assigned to the bacterial concentrations in the sample.  This is due to the fact that bacteria are alive and capable of mortality and reproduction.  Instead of a specific value, an approximate target value of the bacterial concentration is assigned to the sample by the laboratory preparing the positive control sample.

For benthic macroinvertebrate analysis, accuracy will be determined by having 20% of the samples re-analyzed and validated to Level 3 by a professional taxonomist.

Instructions for Determining accuracy (chemical analyses):

Record all results from the test.  Determine the mean value.  Compare the average value to the true value.  Compare this difference to the accuracy objective set in the previous tables.  If the absolute difference is greater, corrective action will be taken to improve performance.  We will consult our technical advisors to determine the appropriate corrective action.

| EXAMPLE: ACCURACY |  |  |  |
| --- | --- | --- | --- |
| During a recent training session, volunteer monitors checked their pH meters against a standard buffer solution of pH 7.0. The following results were read: |  |  |  |
| 7.5 7.4 6.7 | 7.2 6.8 7.3 | 6.5 7.2 6.8 | 7.0 7.4 7.2 |
| Determine the mean result. Most calculators will determine a mean. To calculate: Mean : ACCURACY = mean value - true value To obtain a percent reading: Divide the ACCURACY BY the true value and multiply by 100. The average of these measurements is equal to 7.08. Since we know that the reference or true value is 7.00, the difference between the mean pH value is off or biased by +0.08 units or 1%. This level of accuracy is within the objective of ± 10 percent. Record these results on your QA Form: Data accuracy, Detection Limit, Precision. |  |  |  |

Table 7.6  Example of QA Form:  Data accuracy

| Parameter/ units | Date | sensitivity | accuracy Objective | Estimated Bias | Meet Objective? Yes or No | Corrective action planned | Date Corrective Action taken |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Temperature o C | 5/21/96 | 0.5 | ±10% | -2% -0.5%* * after correction factor given. | Yes | One thermometer was way off, it was discarded. All other thermometers were given a correction factor to improve their accuracy | 5/21/96 |
| Dissolved oxygen (mg/l) | 5/21/96 | 0.2 | sodium thiosulfate 20.00±0.2ml | 21.0 | No | replace reagent | 6/15/96 |
| pH standard units | 5/21/96 | 0.1 | ±10% | -5% | yes | none needed |  |
| conductivity (µS/cm) | 5/21/96 | 100 | ±10% | +10% | yes | none needed |  |
| turbidity (JTU) | 5/21/96 | 5 | ± 5 | +1.4 | yes | none needed |  |

## 7. 2.	Standardization of Instruments and Test Procedures (chemical and physical parameters)

The temperature measurements will be standardized by comparing our thermometers to a NIST-certified or calibrated thermometer.  Instructions for checking the sodium thiosulfate are included in the test kit.  (Additional reagents and glassware must be purchased separately however.)  If the result is unsatisfactory, as indicated in the instructions, the sodium thiosulfate and/or other reagent will be discarded and replaced with new reagents.  The accuracy of the dissolved oxygen test will also be assured by taking these steps:

- Care is taken not to aerate water samples during collection,

- Water is added gently to the dissolved oxygen bottle,

- No air bubbles are present in the sample,

- The titration sample will be measured carefully with a graduated cylinder,

- The sample is swirled thoroughly after each drop of titrant,

- If the endpoint is overrun, another 20 ml. of the sample will be titrated.

Comparators, colorimeters or spectrophotometers and associated reagents will likewise be evaluated twice a year using standards and by comparison to meters at an outside laboratory.

## 7. 3.	Comparability

Comparability is the degree to which data can be compared directly to similar studies.  Citizen monitoring groups will use the following methods to ensure that their data can be compared to others:

- San Francisco Estuary Institute’s Citizen monitoring Protocols,

- U.S. EPA’s Volunteer Monitoring Manuals for Streams, Lakes or Estuaries,

- Yuba River Monitoring Program Volunteer Manual, and

- California’s Department of Fish and Game’s (CDFG) Stream Bioassessment Protocol for Citizen Monitors.

Before modifying these methods, or developing alternative or additional methods, technical advisors will evaluate and review the effects of the potential modification.  It will be important to address their concerns about data quality before proceeding with the monitoring program.

## 7. 4.	Completeness

Completeness is the fraction of planned data that must be collected in order to fulfill the statistical criteria of the project.  Volunteer data will not be used for legal or compliance uses.  There are no statistical criteria that require a certain percentage of data.  However, it is expected that 80% of all measurements could be taken when anticipated.  This accounts for adverse weather conditions, safety concerns, and equipment problems.

We will determine completeness by comparing the number of measurements we planned to collect compared to the number of measurements we actually collected that were also deemed valid.  An invalid measurement would be one that does not meet the sampling methods requirements and the data quality objectives. Completeness results will be checked quarterly.  This will allow us to identify and correct problems.  Table 7.7 will be used to record our completeness information.

Instructions for Determining Completeness:

To determine the percent completed:  Divide the number of valid samples collected and analyzed by the number of samples anticipated in the monitoring design.  multiply by 100%.  In the example below, the volunteers met their objective of 80% completeness for temperature, but not dissolved oxygen.  The volunteers reviewed their sampling methods and realized that some volunteers were not fixing the dissolved oxygen samples correctly.  When they corrected this activity their completeness improved.

Table 7.7  Example QA Form:  Completeness

| Parameter | Collection Period | No. of Samples Anticipated | No. Valid Samples Collected and analyzed | Percent Completed | Comments |
| --- | --- | --- | --- | --- | --- |
| Temperature | 6/1/96 - 9/1/96 | 35 | 33 | 94.3% |  |
| Dissolved oxygen | 6/1/96 - 9/1/96 | 35 | 27 | 77.1% | Volunteers were not fixing samples correctly in field. |
| Temperature | 9/1/96 - 12/1/96 | 35 | 32 | 91.4% |  |
| Dissolved oxygen | 9/1/96 - 12/1/96 | 35 | 32 | 91.4% |  |

## 7. 5.	Precision

Precision describes how well repeated measurements agree.  The precision objectives described here refer to repeated measurements taken by different, trained volunteers or the same volunteer on the same water sample.   Additional variability would be expected if comparisons were made between different samples taken at the same location.

These precision objectives apply to duplicate and split samples taken as part of the QC session or as part of periodic in-field QC checks.  For most parameters, measurements on the same sample read by different volunteers using the same equipment should be within 10% of each other. For turbidity measured in JTU’s, most natural readings should be low (e.g. < 20 JTUs). Because the sensitivity of the method is 5 JTUs, accuracy and precision could be as poor as 100%.  We will compensate for this by conducting actual suspended sediment measurements in addition to turbidity measurements.

Precision for bacterial parameters will be determined by having the same analyst complete the IDEXX procedure for two or more replicates of the same sample. At a minimum this should be done once for every 20 samples.  The results of the replicates should be within ½ of an order of magnitude of each other.

For benthic macroinvertebrate analysis, precision will be determined by having the technical advisor perform an evaluation on the citizen analysts as discussed in Section 14.2 of this QAPP.

Instructions for Determining Precision (chemical analyses):

All volunteers run tests on the same sample.  Record all results from the test.  Determine the mean value.  Calculate the standard deviation.  Compare the standard deviation to the precision objective set in Tables 7.1- 7.4.  If the standard deviation is greater, corrective action will be taken to improve performance.  We will consult our technical advisors to determine the appropriate corrective action.

| EXAMPLE: PRECISION |  |  |  |
| --- | --- | --- | --- |
| During a recent training session, volunteer monitors checked their pH meters against a standard buffer solution of pH 7.0. The following results were read: |  |  |  |
| 7.5 7.4 6.7 | 7.2 6.8 7.3 | 6.5 7.2 6.8 | 7.0 7.4 7.2 |
| Determine the mean result. Most calculators will determine a standard deviation. To calculate: standard deviation: where n is the number of samples, x is the sample value and Precision = standard deviation ) mean H 100. The standard deviation of these measurements is 0.32. The mean is 7.08. The precision is 4.5%. This level of precision is within the objective of 10 percent. Record these results on Your QA Form: Data accuracy, Detection Limit, Precision. |  |  |  |

Table 7.8  Example:  Data Quality Form:  Precision

| Parameter/ units | Date | Mean (x) | Standard Deviation (s.d.) | s.d./x (%) | Precision Objective | Meet Objective? yes or No | Corrective action planned | Date Corrective Action taken |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Temperature o C | 5/21/96 | 22.0 | 0.53 | 2.4% | ± 10% | yes | Precision calculated on corrected thermometers (see accuracy info). |  |
| Dissolved oxygen (mg/l) | 5/21/96 | 8.4 | 1.0 | 11.9% | ± 10% | No | Re-training needed. Volunteers were over-titrating, need to stir thoroughly and add drop by drop near endpoint. | 6/15/96 |
| pH standard units | 5/21/96 | 7.8 | 0.39 | 5.0% | ± 10% | yes | none needed |  |
| conductivity (umhos/cm) | 5/21/96 | 735 | 59 | 8.0% | ± 10% | yes | none needed |  |
| turbidity (JTU) | 5/21/96 | 5 | 2.4 | na | ± 5 JTUs | yes | none needed |  |

## 7. 6.	Representativeness

Representativeness describes how relevant the data are to the actual environmental condition.  Problems can occur if:

- Samples are taken in a stream reach that does not describe the area of interest (e.g. a headwaters sample should not be taken downstream of a point source),

- Samples are taken in an unusual habitat type (e.g. a stagnant backwater instead of in the flowing portion of the creek),

- Samples are not analyzed or processed appropriately, causing conditions in the sample to change (e.g. water chemistry measurements are not taken immediately).

Representativeness will be ensured by processing the samples in accordance with Section 10, 11 and 12, by following the established methods, and by obtaining approval of this document.

## 7. 7.	Sensitivity

The method detection limit is the lowest possible concentration the instrument or equipment can detect.  This is important to record because we can never determine that a pollutant was not present, only that we could not detect it. Sensitivity is the ability of the instrument to detect one concentration from the next.  Sensitivities are noted in Tables 7.1. - 7.5.

# 8.	Training Requirements and Certification

All citizen monitoring leaders must participate in three hands-on training sessions on water quality monitoring conducted by the SYRCL, UC Davis, and the SWRCB.  The following topics are covered under this training:

- General hydrology

- Ecology

- Safety

- Quality Assurance and Quality Control Measures

- Sampling Procedures

- Field Analytical Techniques

- Data recording.

For macroinvertebrate bioassessment citizen monitoring leaders must also participate in a three day training course provided by the California Department of Fish and Game, the Sustainable Lands Stewardship Institute, the American Fisheries Society, or the State Water Resources Control Board.

Trained citizen monitoring leaders may then train their rank-and-file volunteers. Individual trainees are evaluated by their performance of analytical and sampling techniques, by comparing their results to known values, and to results obtained by trainers and other trainees.

In addition to completion of the above described training course, the citizen monitoring leaders must participate in semi-annual quality control sessions conducted by through the Yuba River Monitoring Program Steering Committee.  The semi-annual quality control sessions will provide an opportunity for citizen monitoring groups to check the accuracy and precision of their equipment as well as of their own testing techniques.  The monitor will bring his/her equipment to the session. The monitor will conduct duplicate tests on all analyses and meet the data quality objectives described in Section 7.  If a monitor does not meet the objectives, the trainer will re-train and re-test the monitor.  If there is insufficient time at the QC session to re-train and re-test monitors, the monitor will be scheduled for an additional training session.  The monitor will be encouraged to discontinue monitoring for the analysis of concern until training is completed.

The quality control trainer will examine kits for completeness of components: date, condition, and supply of reagents, and whether the equipment is in good repair.  The trainer will check data quality by testing equipment against blind standards.  The trainer will also ensure that monitors are reading instruments and recording results correctly.  Sampling and safety techniques will also be evaluated.  The trainer will discuss corrective action with the volunteers, and the date by which the action will be taken.  The citizen monitoring leader is responsible for reporting back that the corrective action has been taken.  Certificates of completion will be provided once all corrective action has been completed.

Trainers are defined as water quality professionals from the U.S. Environmental Protection Agency, the State Water Resources Control Board, and the Regional Water Quality Control Boards.  Additional qualified trainers will be recruited and designated by these agencies from experienced citizen monitoring organizations, universities and colleges, commercial analytical laboratories, and  other federal, state, and local agencies.

Please see Section 6.2 Project Timetable for complete list of trainings to be held.  SYRCL staff will be responsible for all components of the Quality Assurance Program Plan and data management.

# 9.	Documentation and Records

All field results will be recorded at the time of completion, using the data sheets (see Appendix 2).  Data sheets will be reviewed for outliers and omissions before leaving the sample site.  Data sheets will be signed after review by the citizen monitoring leader.  Data sheets will be stored in hard copy form at a specified location unique to each citizen monitoring group.  Field sheets are archived for three years from the time they were collected.

If data entry is performed at another location, duplicate data sheets will be used, with the originals remaining at the headquarters site.  Hard copies of all data as well as computer back-up disks are maintained at headquarters.

A maintenance log will also be kept by each citizen monitoring group. This log details the dates of equipment inspection and calibrations, as well as the dates reagents are replaced.

# 10.	Sampling Process Design

## 10.1.	Rationale for Selection of Sampling Sites

Sampling sites are indicated on the chart in Appendix 2.  The following criteria were evaluated when choosing sampling locations:

- access is safe,

- permission to cross private property is granted,

- sample can be taken in main river current or where homogeneous mixing of water occurs,

- sample is representative of the part of the river or beach of interest,

- location complements or supplements historical data,

- location represents an area that possesses unique value for fish and wildlife or recreational use.

If the monitoring program requires reference sites these locations are chosen upstream of any potential impact.  A site chosen to reflect the impact of a particular discharge, tributary or land use should be located downstream of the impact where the impact is completely integrated with the water, but upstream of any secondary discharge or disturbance.

Volunteers are instructed to work in teams of at least two people.  If a scheduled team cannot conduct the sampling together, the available team member will call an additional member.

Prior to final site selection, permission to access the stream is obtained from all property owners.  If access to the site is a problem, the citizen monitoring leader will select a new site.  Safety issues are included in Monitoring Manual.

Sample sites will be reviewed by the leader.  A short report will be made about the site.  The report will describe conditions and include photographs.

## 10.2.	Sample Design Logistics

Volunteers are instructed to work in teams of at least two people.  If a scheduled team cannot conduct the sampling together, the team captain is instructed to contact the citizen monitoring leader so that arrangements can be made for a substitute trained volunteer.

Prior to final site selection, permission to access the stream is obtained from all property owners.  If access to the site is a problem, the citizen monitoring leader will select a new site following the site selection criteria identified in Section 10.1.

Safety measures will be discussed with all volunteers.  No  instream sampling will be conducted if there are small creek flood warnings or advisories. It is the responsibility of the citizen monitoring organization to ensure the safety of their volunteer monitors.  Safety issues are included in the Yuba River Monitoring Program Volunteer Manual .

- Sampling Method Requirements

The Yuba River Monitoring Program Volunteer Monitoring Manual describes the appropriate sampling procedure for collecting samples for water chemistry.  All parameters will be measured with approved protocols from the U.S. EPA Volunteer Stream Monitoring Manual with the exception of suspended sediment.

The suspended sediment protocol will be as follows:

Water sample will be gathered using U.S. EPA protocols.  Using a 1 L wide mouth nalgene bottle – agitate sample and pour directly into 60 mL syringe barrel affixed to sampling device.  The filtering device will contain a pre-weighed 47 mm diameter 1m pore sized glass fiber filter.  The water will be passed through the filter and the filter removed from the filtering device and placed in a pre-weighed aluminum weighing tray.  The filters will be dried at 105 C and weighed.  The weight of the filter and/or the filtering tray will be subtracted from the total weight of the sample to document the weight of the sediment alone.  This value will be used to calculate the concentration of suspended sediment in the original water sample.

At 4 locations in the Yuba basin, the results of this sampling protocol will be compared to the results of the U.S. Geological Survey suspended sediment sampling protocol, which using integrated water sampling and river cross-sectional sampling to determine suspended sediment concentration and load.

Protocol taken from:

Stream Hydrology – an Introduction for Ecologists. Nancy D. Gordon, Thomas A. McMahon, Brian L. Finlayson, John Wiley and Sons, Ltd. Chichester, England.  1992.  526 pp.

All samples will be taken with either a Van Dorn, Niskin, or Kemmerer sampling device, a LaMotte dissolved oxygen sampling device, or by dipping a plastic container into the midstream of a wadeable creek.

Sampling devices will be rinsed three times with sample water prior to taking each sample.   Whenever possible, the collector will sample from a bridge so that the creek is not disturbed from wading.  All samples are taken in mid-stream, at least one inch below the surface.  If it is necessary to wade into the water, the sample collector stands downstream of the sample, taking a sample upstream.  If the collector disturbs sediment when wading, the collector will wait until the effect of disturbance is no longer present before taking the sample.

The following table describes the sampling equipment, sample holding container, sample preservation method and maximum holding time for each parameter.

Table 11.1 Sampling Method Requirements

| Parameter | Sampling Equipment | Preferred / Maximum Holding Times |
| --- | --- | --- |
| Conventional Parameters |  |  |
| Temperature | plastic or glass container or sample directly | immediately |
| Dissolved oxygen | glass D.O. bottle | immediately / fix per protocol instructions, continue analysis within 8 hr. |
| pH | plastic or glass container | immediately |
| conductivity | plastic or glass container | immediately / refrigerate up to 28 days |
| turbidity | plastic or glass container | immediately / store in dark for up to 24 hr. |
| Nutrients |  |  |
| Ammonia | Van Dorn, LaMotte or plastic sampling bottle | immediately |
| Nitrates | Van Dorn, LaMotte or plastic sampling bottle | immediately / refrigerate in dark for up to 48 hr. |
| Phosphate | Van Dorn, LaMotte or plastic sampling bottle | immediately |
| Urban Pollutants – Field Measurements |  |  |
| Mercury | Van Dorn, LaMotte or plastic sampling bottle | immediately |
| Arsenic | Van Dorn, LaMotte or plastic sampling bottle | immediately |
| Total Copper | Van Dorn, LaMotte or plastic sampling bottle | immediately |
| Zinc | Van Dorn, LaMotte or plastic sampling bottle | immediately |
| Laboratory Analysis of Chemical Parameters |  |  |
| Total Organic Carbon | acid and d.i. water rinsed glass sampling bottle, teflon liner in lid | refrigerate to 4 degrees C, send to lab immediately |
| Metals | plastic sampling bottle | fix with Ultrapure (or comparable) nitric acid, send to lab immediately |
| Oil and Grease | acid and d.i. water rinsed glass sampling bottle, teflon liner in lid | refrigerate to 4 degrees C, send to lab immediately |
| PAH’s | acid and d.i. water rinsed glass sampling bottle, teflon liner in lid | refrigerate to 4 degrees C, send to lab immediately |
| Pesticides and other synthetic organic compounds | acid and d.i. water rinsed glass sampling bottle, teflon liner in lid | refrigerate to 4 degrees C, send to lab immediately |
| Toxicity | acid and d.i. water rinsed glass sampling bottle, teflon liner in lid | refrigerate to 4 degrees C, send to lab immediately |
| Biological Samples |  |  |
| Bacteria | Sterile plastic sampling bottle or whirl-pack | Refrigerate in the dark; deliverd to the lab within 4 hours, start analysis within 6 hours |
| Benthic macroinvertebrates | Glass bottles | Fixed with ethanol immediately |

# 12.	Sample Handling and Custody Procedures

## 12.1.	Sample Handling

Identification information for each sample will be recorded on the field data sheets (see Appendix 2) when the sample is collected.  Samples are normally processed in the field.  Split samples and samples that are not processed immediately will be labeled with the waterbody name, sample location, sample number, data and time of collection, sampler’s name, and method used to preserve sample (if any).

## 12.2.	Custody Procedures

The conventional water quality monitoring tests do not require specific custody procedures since they will, in most cases, be conducted immediately by the same person who performs the sampling.  In certain circumstances (such as driving rain or extreme cold), samples will be taken to a nearby residence for analysis.  The dissolved oxygen samples will be fixed prior to transport.

When samples are transferred from one volunteer to another member of the citizen monitoring group for analysis, or from the citizen monitoring program to an outside professional laboratory, then a Chain of Custody form should be used.  This form identifies the waterbody name, sample location, sample number, data and time of collection, sampler’s name, and method used to preserve sample (if any).  It also indicates the date and time of transfer, and the name and signature of the sampler and the sample recipient. It is recommended that the Chain of Custody form used be the one provided by the outside professional laboratory.  When quality control checks are performed by a professional lab, their samples will be processed under their chain of custody procedures with their labels and documentation procedures.

For benthic macroinvertebrate samples, the California Department of Fish and Game Aquatic Bioassessment Laboratory Chain of Custody form will be used.

## 12.3.	Disposal

All analyzed samples (except for waste from the nitrate/cadmium reduction test and the Nessler ammonia test) including used reagents, buffers or standards will be collected in a plastic bottle clearly marked “Waste” or “Poison”.  This waste material will be disposed of according to appropriate state and local regulations.  This will usually mean disposal into a drain connected to a sewage treatment plant.

Liquid waste from the cadmium reduction nitrate test will be kept separate and disposed of at a facility that is permitted to handle, transport, or dispose Cd waste. Liquid waste from the Nessler ammonia test (which contains mercury) will likewise be kept separate and disposed of at a facility that is permitted to handle, transport, or dispose Hg waste.  Waste from the zinc reduction nitrate test and the salicylate ammonia test can be held in the regular waste container and disposed of as described in the previous paragraph.

Whenever possible, if waste includes reagents from the detergent test, these wastes will be poured down a drain underneath a flume hood.

# 13.	Analytical Methods Requirements

Water chemistry is monitored using protocols outlined in the Yuba River Monitoring Program Volunteer Manual.  The methods were chosen based on the following criteria:

- capability of volunteers to use methods,

- provide data of known quality,

- ease of use,

- methods can be compared to professional methods in Standard Methods.

If modifications of methods are needed, comparability will be determined by side-by-side comparisons with a US EPA or APHA Standard Method on no less than 50 samples.  If the results meet the same precision and accuracy requirements as the approved method, the new method will be accepted.

Table 13.1 outlines the methods to be used, any modifications to those methods, and the appropriate reference to a standard method.

Table 13.1  Analytical Methods for Water Quality Parameters

| Parameter | Method | Modification | Reference (a) |
| --- | --- | --- | --- |
| Temperature | Thermometric | Alcohol-filled thermometer marked in 0.5oC increments | 2550 B. |
| Dissolved Oxygen | Winkler Method, Azide Modification | Prepackaged reagents, 20 ml sample size | 4500-O C. |
| Dissolved Oxygen | Membrane Electrode | none | 4500-O G. |
| pH | Electrometric | none | 4500-H B. |
| Conductivity | Electrometric | none | 2520 B. |
| Turbidity | Nephelometer | none | None |
| Ammonia | Nessler | prepackaged reagents, colorimeter or spectrophotometer | 4500 – NH3 C 18th edition only (1992) |
| Nitrate | Cadmium Reduction | prepackaged reagents, colorimeter or spectrophotometer | 4500 – NO3- E. |
| Ortho-Phosphate | Ascorbic acid | prepackaged reagents, colorimeter or spectrophotometer | 4500 – P E. |
| Total Copper | Neocuproine | Color Comparator | 3500 -Cu D. |
| Fecal Coliform Bacteria | Colilert 18 hour | none | 9223 |
| Benthic Macroinvertebrates | California Stream Bioassessment Protocol | Level 2 (to family only) | Harrington, Jim, CDFG, 1997 |

- All of the above cited methods, with the exception of turbidity and benthic macroinvertebrates are described in Standard Methods for the Examination of Water and Wastewater:

Andrew D. Eaton, Lenore S. Clesceri, Arnold E. Greenberg, Mary Ann H. Franson.

Standard Methods for the Examination of Water and Wastewater, prepared and published jointly By American Public Health Association, American Water Works Association, Water Environment Federation, 20th edition, Washington, DC : American Public Health Association, 1998.

# 14.	Quality Control Requirements

Quality control samples will be taken to ensure valid data are collected.  Depending on the parameter, quality control samples will consist of field blanks, replicate samples, or split samples.  In addition, quality control sessions (a.k.a. intercalibration exercises) will be held twice a year to verify the proper working order of equipment, refresh volunteers in monitoring techniques and determine whether the data quality objectives are being met.

## 14.1.	Cautions Regarding  Test Procedures

### 14.1.1.	Dissolved oxygen Test

The Winkler method is not appropriate for highly alkaline waters.

Other citizen monitoring groups have noted problems with short shelf-life of the sodium thiosulfate reagent.  Field measurements should be evaluated immediately to determine whether they are reasonable.

### 14.1.2.	Nutrients

When mixing nitrate reagents take care not to agitate aggressively.  The LaMotte phosphate reagents have been shown to degrade well within their listed shelf life once opened.

### 14.1.3.	Urban Pollutants

Urban street surface contaminants (e.g. sediment, oil and grease) and fecal coliform will be measured.

The low sensitivity of the copper test may preclude detecting copper as most falls out of solution forming copper carbonate.

## 14.2.	Field/Lab Blanks, Duplicate Field Samples, and Split Samples

Table 14.1 describes the quality control regimen.

Field/Laboratory Blanks: For turbidity and specific chemical analysis (see Table 14.1) performed in the field Field blanks (a.k.a. reagent blanks) will be taken once every 20 samples, or quarterly whichever comes first except for nutrient sampling. For nutrients using comparators, a reagent blank sample will be analyzed every sampling trip.  Color can sometimes appear in these nutrient blanks, suggesting that the real samples may be overestimating the true nutrient concentration. When colorimeters or spectrophotometers are used at the group’s facility for nutrient analysis, a laboratory reagent blank will be analyzed and recorded for each day of analysis.  For bacterial analysis performed at a group’s facility a laboratory blank will be performed during for each sampling/analysis event.

Instructions for Field and Lab  Blanks:  Distilled water is taken into the field or used in the laboratory and handled just like a sample.  It will be poured into the sample container and then analyzed.  Field blanks are recorded on the normal sampling datasheet.  For nutrients measured with comparators, results from the field blanks should be “not detected”.  If nutrients are detected, corrective action will be taken to eliminate the problem.  For nutrients measured with colorimeters, the reagent blanks should be less than 0.05 ppm and the specific value should be recorded and subtractedfrom the field sample result. For bacterial analysis, the reagents are added to distilled water (in the same manner as for a field sample) and that “blank” is then sealed in a quantitray and incubated along with the field samples.  The blank should be below detection limits at the end of the incubation period.

Duplicate Field Samples:  For chemical, physical, and bacterial analysis duplicate field samples will be taken once every 20 samples, or quarterly whichever comes first.  Duplicate samples will be collected as soon as possible after the initial sample has been collected, and will be subjected to identical handling and analysis.

For benthic macroinvertebrate sampling, instead of duplicate sampling, each sampler will be evaluated annually by measuring the area sampled upstream of the net.  The area should be two square feet and should be verified by using a two square foot pvc frame.

Split Samples: Twice a Year, split spiked samples (standards) will be analyzed as part of the Quality Control (Intercalibration) Session.  The split standard is one sample, containing a known concentration of an analyte, that is divided equally into two or more sample containers.  Split standards will be analyzed by the volunteers, and sent to a professional laboratory (except for dissolved oxygen, temperature, and pH), before the maximum sample handling time is exceeded.  Volunteers will analyze the split standard normally and will perform at least three analyses on that same sample. From these results accuracy and precision will be determined.  The professional laboratory will analyze the sample using the method referenced in Table 13.1

For turbidity, the laboratory will use the nephelometric method, even though these results are not strictly comparable to the visual JTU comparators.  The results of turbidity using the two methods will be plotted to determine if there is a linear correlation.  If this correlation is significant, then it will be used to estimate and compare results of the turbidity tubes with nephlometric results.  The Technical Advisory Committee will use the product-moment correlation coefficient (r) to determine the adequacy of the correlation.

For bacteria, split field samples or split positive controls will be analyzed by an outside professional laboratory twice annually.

A minimum 20% of the benthic macroinvertebrate samples will be subjected to validation by a outside professional taxonomist.  Following analysis by the citizen group the selected samples will be reconstituted and sent out for professional level 3 taxonomic analyses. Reconstituted means opening the vials containing the 100 identified specimens, pouring the specimens back into the original sample jar, and gently stirring the contents.  In addition, once a year citizen macroinvertebrate analysts will participate in an intercalibration exercise in which their subsampling/sorting and taxonomic skills will be evaluated.  A minimum of two teams of analysts will each inspect each other’s processed grids immediately following completion of the subsampling procedure.  There should be no more than 10% missed organisms.  A technical advisor should then evaluate each of the citizen analysts by testing their identification to order and family level on at least 20 specimens, including at least one representative from each of the major orders and families as determined by the technical advisor for that watershed. Accuracy and precision can be determined by the results of these validation and evaluation measures.

Table 14.1  Quality Control Requirements

| Parameter | Blank | Duplicate Sample | Split Sample to lab | QC session (intercal.) |
| --- | --- | --- | --- | --- |
| Water quality |  |  |  |  |
| Temperature | none | 5% or a minimum of once a year | none | twice a year |
| Dissolved oxygen | none | 5% or a minimum of once a year | none | twice a year |
| pH | none | 5% or a minimum of once a year | none | twice a year |
| conductivity | 5% | 5% or a minimum of once a year | twice a year | twice a year |
| turbidity | 5% | 5% or a minimum of once a year | twice a year | twice a year |
| Nutrients (comparators) |  |  |  |  |
| Ammonia | daily | 5% or a minimum of once a year | twice a year | twice a year |
| Nitrate | daily | 5% or a minimum of once a year | twice a year | twice a year |
| Phosphate | daily | 5% or a minimum of once a year | twice a year | twice a year |
| Nutrients (colorimeters or spectrophotometers) |  |  |  |  |
| Ammonia | daily | 5% or a minimum of once a year | twice a year | twice a year |
| Nitrate | daily | 5% or a minimum of once a year | twice a year | twice a year |
| Phosphate | daily | 5% or a minimum of once a year | twice a year | twice a year |
| Urban Pollutants |  |  |  |  |
| Suspended Sediment | 5% | 5% or a minimum of once a year | twice a year | twice a year |
| Hydrocarbons | 5% | 5% or a minimum of once a year | twice a year | twice a year |
| Salts | 5% | 5% or a minimum of once a year | twice a year | twice a year |
| Metals | 5% | 5% or a minimum of once a year | twice a year | twice a year |
| Biological Parameters |  |  |  |  |
| Benthic Invertebrates | none | None, instead conduct evaluation of sampling area annually | 20% per year | once a year |
| Total & Fecal Coliform | daily | 5% or a minimum of once a year | twice a year | twice a year |

# 15.	Instrument/Equipment Testing, Inspection and Maintenance Requirements

A maintenance log is kept by the monitoring group leader.  This log records reagent use, and any problems noted with equipment.  Calibration information is recorded on the datasheets.

## 15.1.	Temperature

Before each use, thermometers are checked for breaks in the column.  If a break is observed, the alcohol thermometer will be placed in nearly boiling water so that the alcohol expands into the expansion chamber, and the alcohol forms a continuous column. verify accuracy by comparing with a calibrated or certified thermometer.

## 15.2.	Dissolved oxygen

Before each use, bottles, droppers, and color comparators are checked to see if they are clean and in good working order.  Reagents are replaced annually according to manufacturer’s recommendation.

## 15.3.	pH and conductivity

Before each use, pH and conductivity meters are checked to see if they are clean and in good working order.  pH and conductivity meters are calibrated before each use.  pH buffers and conductivity standards are replaced at least annually.  conductivity standards are stored with the cap firmly in place and in a dry place kept away from extreme heat.  Do not re-use pH or conductivity standards.

## 15.4.	turbidity

Before each use, turbidity tubes are checked to ensure that they are clean.  The turbidity standard will be replaced annually.

## 15.5.	Nutrients and Urban Pollutants

Before each use, test kits are checked to ensure that droppers, sample containers, and color comparators are clean and in working condition.  Reagents are replaced annually according to manufacturer’s instructions.

16.	Instrument Calibration and Frequency (chemical and physical parameters)

Instruments will be calibrated accordingly to the following schedule.  Standards will be purchased from a chemical supply company or prepared by a laboratory certified by U.S. EPA for chemical analysis of  water or wastewater.  All meters (pH, conductivity, oxygen) have been calibrated and will be evaluated monthly prior to use with known  standards and annually by comparison to meters at an outside laboratory. The pH meters will be calibrated using low ionic strength known pH buffers.  The conductivity meters will be calibrated using mixed salt solution of known composition and conductivity.  The oxygen meters will be calibrated using “0” oxygen solution and a solution saturated with oxygen (100%).  The ammonia/nitrogen and the nitrate/nitrogen test kits will be standardized with solutions of known dissolved nitrogen concentration.  The suspended sediment protocol will be standardized by taking a known (weighed) amount of sediment and suspending it in a known volume of water in the sampling container.  Subsamples will be taken from the sampling container and the sediment will be filtered (removed) from the water and weighed.  The observed concentration (mg/L) of suspended sediment will be compared to the expected concentration based on the known amount of sediment added to the known volume of water.  If necessary, the observed concentrations will be corrected based on this calibration.  The dissolved oxygen kit (Winkler method) will be checked by monthly standardization and/or replace of the sodium thiosulfate solution in the test kit, and/or by comparing the entire kit to a saturated oxygen standard.  Calibration records will be kept at a location where they can be easily accessed before and after equipment use.  This will likely be at the citizen monitoring organization’s main office or the volunteer monitor’s home.

Table 16.1  Instrument Calibration and frequency

Conventional Water quality Parameters

| Equipment Type | Calibration Frequency | Standard or Calibration Instrument Used |
| --- | --- | --- |
| Temperature | Every 6 months | NIST calibrated or certified thermometer |
| Dissolved Oxygen (Winkler) | check sodium thiosulfate every sampling day | titration |
| Dissolved Oxygen meter | Every sampling day | At a minimum, water saturated air, according to manufacturer’s instructions. |
| pH | Every sampling day | pH 7.0 buffer and pH 3.0 buffer |
| conductivity | Every sampling day | conductivity standard |
| Turbidity meter (nephelometer) | Every sampling day | For clear ambient conditions use an 1.0 NTU standard, for turbid conditions use an 10.0 NTU standard |

Nutrients (using comparators)

| Equipment type | Calibration frequency | Standard or Calibration Instrument Used |
| --- | --- | --- |
| Ammonia | every 6 months or when reagents replaced | ammonia standard |
| Nitrate | every 6 months or when reagents replaced | nitrate standard |

Nutrients (using colorimeters or spectrophotometers)

| Equipment type | Calibration frequency (test standard) | Standard or Calibration Instrument Used |
| --- | --- | --- |
| Ammonia | Every day of analysis | ammonia standard |
| Nitrate | Every day of analysis | nitrate standard |

Urban Pollutants

| Equipment type | Calibration frequency | Standard or Calibration Instrument Used |
| --- | --- | --- |
| Suspended Sediment | every 6 months | standardize with known concentrations |
| Oil and Grease | every 6 months | Hydrocarbon standard |
| De-icing Salts | every 6 months | Salinity standard |

# 17.	Inspection/Acceptance Requirements

Upon receipt, buffer solutions, standards, and reagents used in the field kits will be inspected by the citizen monitoring leader for leaks or broken seals, and to compare the age of each reagent to the manufacturer’s recommended shelf-life.  All other sampling equipment will be inspected for broken or missing parts, and will be tested to ensure proper operation.

Before usage, thermometers are inspected for breaks.  Breaks can be eliminated by heating (see Section 15.1).  If not, they will be returned to the manufacturer.

Reagents are replaced before they exceed manufacturer’s recommended shelf life.  These shelf lives are typically one to two years.  However, specific replacement dates can determined by providing the reagent lot number to the LaMotte company by phone at (800) 344-3100 or facsimile at (410) 778-6394.  Reagent replacement dates are noted in the maintenance log.

# 18.	Data Acquisition Requirements

## 18.1.	Analytical Data

Only certified analytical laboratories or academic laboratories (with approval of State and/or Regional Board staff) will be used for quality assurance checks.  The Technical advisory Committee (TAC) or technical advisors will review these laboratories’ data as well as the volunteers.  They will review the lab’s own quality control data to ensure data validaty.

## 18.2.	Geographical Information/ Mapping

USGS maps will be used to verify watershed boundaries and river courses.  NOAA navigation charts can be marine used for mapping marine sampling sites.  Additional information on distribution of natural resources will be obtained from the National Park Service and the CDFG’s Biodiversity database.  Land use information will be obtained from local planning offices.  When information is requested, the agency will be asked to provide appropriate megadata and any information on data limitations.  This information will be maintained with the data files.

# 19.	Data Management

Field data sheets are checked and signed in the field by the citizen monitoring leader.  The citizen monitoring leader will discard any results where holding times have been exceeded, sample identification information is incorrect, samples were inappropriately handled, or calibration information is missing or inadequate.

Independent laboratories will report their results to the citizen monitoring leader.  The leader will verify sample identification information, review the chain-of-custody forms, and identify the data appropriately in the database.  These data are also reviewed by the technical advisors quarterly.

The data management coordinator will review the field sheets and enter the data deemed acceptable by the citizen monitoring leader and the technical advisors.  Data will be entered into a spreadsheet or a database using a format that is compatible with the Regional WQCB’s database guidelines.  The data coordinator will review electronic data, compare to the original data sheets and correct entry errors.  After performing data checks, and ensuring that data quality objectives have been met, data analysis will be performed.

Raw data will be provided to the State WRCB and Regional WQCB in electronic form at least once every two years so that it can be included in the 305(b) report.  Appropriate quality assurance information can be provided upon request.

# 20.	Assessment and Response Actions

Review of all field and data activities is the responsibility of the citizen monitoring leader, with the assistance of the technical advisory committee.  Volunteers will be accompanied by the citizen monitoring leader, or a technical advisor on at least one of their first 5 sampling trips.  If possible, volunteers in need of performance improvement will be retrained on-site.  All volunteers must attend a refresher course offered by the Yuba River Monitoring Program.  If errors in sampling technique are consistently identified, retraining may be scheduled more frequently.

Within the first three months of the monitoring project, the State Water Board staff, or its designee, will evaluate field and laboratory performance and provide a report to the citizen monitoring group.  All field and laboratory activities, and records may be reviewed by state and EPA quality assurance officers as requested.

# 21.	Reports

The technical advisors will review draft reports to ensure the accuracy of data analysis and data interpretation.  Raw data will be made available to data users per their request.  The individual citizen monitoring organizations will report their data to their constituents after quality assurance has been reviewed and approved by their technical advisors.  Every effort will be made to submit data and/or a report to the State and/or Regional Board staff in a fashion timely for their data uses, e.g. 305(b) report or special watershed reports.

# 22.	Data Review, Validation and Verification

Data sheets or data files are reviewed quarterly by the technical advisors to determine if the data meet the Quality Assurance Project Plan objectives.  They will identify outliers, spurious results or omissions to the citizen monitoring leader.  They will also evaluate compliance with the data quality objectives.  They will suggest corrective action that will be implemented by the citizen monitoring leader.  Problems with data quality and corrective action will be reported in final reports.

# 23.	Validation and Verification Methods

As part of standard field protocols, Any sample readings out of the expected range will be reported to the citizen monitoring leader.  A second sample will be taken as soon as possible to verify the condition.  It is the responsibility of the citizen monitoring leader to re-train volunteers until performance is acceptable.

# 24.	Reconciliation with DQOs

The Technical Advisory Committee will review data quarterly to determine if the data quality objectives (DQOs) have been met.  They will suggest corrective action.  If data do not meet the project’s specifications, the following actions will be taken.  First, the technical advisors will review the errors and determine if the problem is equipment failure, calibration/maintenance techniques, or monitoring/sampling techniques.  If the problem cannot be corrected by training, revision of techniques, or replacement of supplies/equipment, then the technical advisors and the TAC will review the DQOs and determine if the DQOs are feasible.  If the specific DQOs are not achievable, they will determine whether the specific DQO can be relaxed, or if the parameter should be eliminated from the monitoring program.  Any revisions to DQOs will be appended to this QA plan with the revision date and the reason for modification.  The appended QA plan will be sent to the quality assurance panel that approved this plan.  When the appended QA plan is approved, the citizen monitoring leader will work with the data coordinator to ensure that all data meeting the new DQOs are entered into the database.  Archived data can also be entered.

Blank form included in Appendix 1.   Data Quality Forms