## QA Table summary
A. QA Tables Summary
Table 1. Summary of MDL study conducted for diazinon
| Sample Description | Nominal | Measured | % RECOVERY | MDL
|  | Concentration | Concentration
|  | (μg/L) | (μg/L)
| MDL .020 μg/LA | 0.02 | 0.021109 | 105.5451
| MDL .020 μg/LA | 0.02 | 0.023889 | 119.4444
| MDL .020 μg/LA | 0.02 | 0.020582 | 102.9078
| MDL .020 μg/LA | 0.02 | 0.02038 | 101.9001
| MDL .020 μg/LA | 0.02 | 0.025918 | 129.5917 | average=
| MDL .020 μg/LB | 0.02 | 0.02 | 100 | 0.02011
| MDL .020 μg/LB | 0.02 | 0.023 | 115 | Standard deviation=
| MDL .020 μg/LB | 0.02 | 0.014 | 70 | 0.006122
| MDL .020 μg/LB | 0.02 | 0.023 | 115 | t= (16, .01)
| MDL .020 μg/LB | 0.02 | 0.022 | 110 | 2.583
| MDL .020 μg/LB | 0.02 | 0.038 | 190 | MDL=
| MDL .020 μg/LB | 0.02 | 0.018 | 90 | 0.015813
| MDL .020 μg/LB | 0.02 | 0.014 | 70
| MDL .020 μg/LB | 0.02 | 0.014 | 70
| MDL .020 μg/LB | 0.02 | 0.013 | 65
| MDL .020 μg/LB | 0.02 | 0.015 | 75
| MDL .020 μg/LB | 0.02 | 0.016 | 80
*Does not include CVs greater than 15%
Superscript denotes when ELISA was conducted A=2/23/01 B=9/10/01
Table 2. Summary of MDL study conducted for chlorpyrifos.
| Sample Description | Nominal | Measured | % RECOVERY | MDL
|  | Concentration | Concentration
|  | (μg/L) | (μg/L) |  | (μg/L)
| MDL .040 μg/LA | 0.04 | 0.04956 | 123.9
| MDL .040 μg/LA | 0.04 | 0.03093 | 77.325
| MDL .040 μg/LA | 0.04 | 0.03915 | 97.875
| MDL .040 μg/LA | 0.04 | 0.03527 | 88.175
| MDL .040 μg/LA | 0.04 | 0.04303 | 107.575
| MDL .040 μg/LA | 0.04 | 0.05001 | 125.025
| MDL .040 μg/LA | 0.04 | 0.04621 | 115.525 | average=
| MDL .040 μg/LA | 0.04 | 0.04499 | 112.475 | 0.044373
| MDL .040 μg/LA | 0.04 | 0.0648 | 162 | Standard deviation=
| MDL .040 μg/LA | 0.04 | 0.05589 | 139.725 | 0.0087523
| MDL .040 μg/LB | 0.04 | 0.053616 | 134.0411 | t= (18, .01)
| MDL .040 μg/LB | 0.04 | 0.042618 | 106.5443 | 2.552
| MDL .040 μg/LB | 0.04 | 0.047643 | 119.1075 | MDL=
| MDL .040 μg/LB | 0.04 | 0.037712 | 94.2812 | 0.022336
| MDL .040 μg/LB | 0.04 | 0.052908 | 132.2689
| MDL .040 μg/LB | 0.04 | 0.03655 | 91.3754
| MDL .040 μg/LB | 0.04 | 0.034792 | 86.9789
| MDL .040 μg/LB | 0.04 | 0.034224 | 85.5609
| MDL .040 μg/LB | 0.04 | 0.043177 | 107.9414
*Does not include CVs greater than 15%
Superscript denotes when ELISA was conducted A=2/23/01 B=8/24/01
Table 3. Summary of precision and accuracy study conducted for diazinon.
| Sample Description | Nominal | Measured | % RECOVERY
|  | Concentration | Concentration
|  | (μg/L) | (μg/L)
| 0.030 μg/L Calibrator | 0.03 | 0.028 | 93.3333
| 0.100 μg/L Calibrator | 0.1 | 0.092 | 92
| 0.250 μg/L Calibrator | 0.25 | 0.264 | 105.6
Study was conducted on 2/21/01
Table 4. Summary of precision and accuracy study conducted for chlorpyrifos.
| Sample Description | Nominal | Measured | % RECOVERY
|  | Concentration | Concentration
|  | (μg/L) | (μg/L)
| .050 μg/L CalibratorA | 0.05 | 0.057315 | 114.6293
| .125 μg/L CalibratorA | 0.125 | 0.14978 | 119.8259
| .300 μg/L CalibratorA | 0.25 | 0.19728 | 78.912
| .050 μg/L CalibratorB | 0.03 | 0.052591 | 175.3028
| .125 μg/L CalibratorB | 0.1 | 0.17748 | 177.4843
| .300 μg/L CalibratorB | 0.25 | 0.36663 | 146.652
| .050 μg/L CalibratorB | 0.05 | 0.055872 | 111.7437
| .125 μg/L CalibratorB | 0.125 | 0.16882 | 135.0548
| .300 μg/L CalibratorB | 0.3 | 0.33145 | 110.4847
Superscript denotes when ELISA was conducted A=2/22/01 B=2/28/01
Table 5. Summary of method blank study conducted for chlorpyrifos.
| Sample Description |  |  | Measured
Concentration
(μg/L)
| Method Blank # 1: Deionized water |  |  | 0.028
| Method Blank # 2: Glass distilled water filtered a .2 nylon filter |  |  | 0.028
| Method Blank # 3: The chlorpyrifos stock solution of 1ppb is filtered through a .2 micron |  |  | 0.026
nylon filter. The filtering apparatus then undergoes the cleaning procedure before the filter
is replaced and glass distilled water is filtered through. This water is tested is method blank 3.
| Method Blank # 4: Glass distilled water filtered through a .2 micron nylon filter using a disposable plastic syringe |  |  | 0.022
| Standard 0 Calibrator |  |  | 0.006
Study was conducted on 2/21/01
Table 6. Summary of method blank study conducted for diazinon.
| Sample Description |  |  | Measured
Concentration
(μg/L)
| Method Blank # 2: Glass distilled water filtered a .2 nylon filter |  |  | 0.016045
| Method Blank # 3: The diazinon stock solution of XXppb is filtered through a .2 micron |  |  | 0.014522
nylon filter. The filtering apparatus then undergoes the cleaning procedure before the filter
is replaced and glass distilled water is filtered through. This water is tested is method blank 3.
| Method Blank # 4: Glass distilled water filtered through a .2 micron nylon filter using a disposable plastic syringe |  |  | 0.015125
| Standard 0 Calibrator |  |  | 0.013
Study was conducted on 2/21/01
Table 7. Summary of filter study-diazinon (conducted to determine the effect of filter type and size on analyte recovery)
| Filter Description: Pore Size/Type | Filtration Method | Nominal Diazinon Concentration | Measured Concentration (ng/L) | % Recovery (Based on Nominal Values) | Conclusions
| .2 µ micron nylon filter | 1 | 0.05 | 0.039 | 79 | Recoveries generally good for all pore sizes and types tested. Unusually high recoveries with.45 unylon filters may be due to contamination.
| .2 µ micron nylon filter | 1 | 0.05 | 0.045 | 90
| .2 µ micron nylon filter | 2 | 0.05 | 0.044 | 88
| .2 µ micron nylon filter | 2 | 0.05 | 0.063 | 126
| .45 µ nylon filter | 1 | 0.05 | 0.079 | 158
| .45 µ nylon filter | 1 | 0.05 | 0.106 | 212
| 1µ glass fiber filter | 1 | 0.05 | 0.05 | 99
| 1µ glass fiber filter | 1 | 0.05 | 0.05 | 100
| .2 µ nylon filter | 1 | 0.5 | 0.456 | 91 | Recoveries generally good for all pore sizes and types tested.
| .2 µ nylon filter | 1 | 0.5 | 0.456 | 91
| .2 µ nylon filter | 2 | 0.5 | 0.468 | 94
| .2 µ nylon filter | 2 | 0.5 | 0.575 | 115
| .45 µ nylon filter | 1 | 0.5 | 0.455 | 91
| .45 µ nylon filter | 1 | 0.5 | 0.423 | 85
| 1µ glass fiber filter | 1 | 0.5 | 0.478 | 96
| 1µ glass fiber filter | 1 | 0.5 | 0.49 | 98
§ Filtration Methods:
o Method 1: pneumatic pump cleaned with rinses of glass distilled water and methanol between samples (cleaning protocol followed, refer to XXX)
o Method 2: disposable syringe and disposable filter attachment
§ Concentrations tested were 50 ng/L and 500 ng/L diazinon spikes; the spike values are 0.029 for .050 µg/L spike C and .268 (at 100%= .536) for the .500 µg/L spike @50% C
§ No reported values have % CVs exceeding 10%. Highlighted cells are values exceeding the lower or upper limits of the linear range for that particular ELISA run.
§ Study was conducted on 2/23/01
Table 8. Filter study -chlorpyrifos
| Filter Description: Pore Size/Type | Filtration Method | Nominal Diazinon Concentration | Measured Concentration (ng/L) | % Recovery (Based on Nominal Values) | Notes
| .2 µ nylon filter | 1 | .050 μg/L | 0.022 | 44 | Trend of increasing recovery with increasing filter pore size. Filter type may or may not have an effect, however this is hard to determine since the filter types tested had different pore sizes. Filtration method 2 seemed to result in decreased recovery –unable to fully determine since these numbers fall below the lower limit of the linear range.
| .2 µ nylon filter | 1 | .050 μg/L | 0.018 | 37
| .2 µ nylon filter | 2 | .050 μg/L | 0.013 | 25
| .2 µ nylon filter | 2 | .050 μg/L | 0.01 | 20
| .45 µ nylon filter | 1 | .050 μg/L | .034 (13%) | 69
| .45 µ nylon filter | 1 | .050 μg/L | 0.026 | 53
| 1µ nylon filter | 1 | .050 μg/L | 0.042 | 83
| 1µ nylon filter | 1 | .050 μg/L | 0.041 | 82
| .2 µ nylon filter | 1 | .500 μg/L | 0.147 | 29 | Trend similar to above. This study would seem to suggest that a significant amount of the analyte, chlorpyrifos, is lost when filtering using the nylon filters with pore sizes ranging form .2 to .45 micron. Recovery is better when filtration is done using the glass fiber filter of 1 micron pore size.
| .2 µ nylon filter | 1 | .500 μg/L | 0.156 | 31
| .2 µ nylon filter | 2 | .500 μg/L | 0.013 | 3
| .2 µ nylon filter | 2 | .500 μg/L | 0.012 | 2
| .45 µ nylon filter | 1 | .500 μg/L | 0.267 | 53
| .45 µ nylon filter | 1 | .500 μg/L | .248 (17%) | 50
| 1µ glass fiber filter | 1 | .500 μg/L | 0.479 | 96
| 1µ glass fiber filter | 1 | .500 μg/L | 0.415 | 83
§ In bold are values that had %CVs exceeding 10%. The CV is noted parenthetically.
§ Filtration Methods:
o Method 1: pneumatic pump cleaned with rinses of glass distilled water and methanol between samples (cleaning protocol followed, refer to XXX)
o Method 2: disposable syringe and disposable filter attachment
§ Concentrations tested were 50 ng/L and 500 ng/L chlorpyrifos spikes.
§ Highlighted cells are values exceeding the lower or upper limits of the linear range for that particular ELISA run.
§ Study was conducted on 2/23/01
Table 9. Summary of results of 50 and 500 pptrillion chlorpyrifos spike
| Nominal Spike Concentration (ng/L) | Measured Concentration (ng/L) | Notes
| .050 ppt@100% | 0.06
| .500ppt@50% | 0.223 | 100%: .446 ppb
§ In bold are values that had %CVs exceeding 10%. The CV is noted parenthetically.
§ Highlighted cells are values exceeding the lower or upper limits of the linear range for that particular ELISA run.
§ Study was conducted on 2/23/01
Table 10. Summary of results of diazinon evaluation sample. The the evaluation sample was made by the laboratory QA officer using ChemService………
| Sample Description | Nominal | Measured | % RECOVERY
|  | Concentration | Concentration
|  | (μg/L) | (μg/L)
| Diazinon Evaluation Sample | 0.25 | 0.25296 | 101%
* ELISA was conducted on 2/21/01
Table 11. Summary of results of Chlorpyrifos evaluation sample. The the evaluation sample was made by the laboratory QA officer using ChemService………
| Sample Description | Nominal | Measured | % RECOVERY
|  | Concentration | Concentration
|  | (μg/L) | (μg/L)
| Diazinon Evaluation Sample | 0.2 | 0.202 | 101%
* ELISA was conducted on 2/28/01

## QA.Total
QA FOR ANALYTE: CHLORPYRIFOS |  | %CV | ng/L
Standard 50 A | Chlor | 4.9781 | 57.3146 | Key
Standard 125 A | Chlor | 10.4837 | 149.7823 | 02/21/2001=A
Standard 300 A | Chlor | 4.5051 | 197.2799 | 02/23/2001=B
Method Blank #2 A | Chlor | 3.2841 | 27.8321 | 02/28/2001=C
Method Blank #3 A | Chlor | 1.5268 | 27.9167
Method Blank #4 A | Chlor | 2.5963 | 26.0541
Method Blank #1 (HPLC H20) A | Chlor | 2.0049 | 21.915
50 ng/l Spike B | Chlor | 5.6275 | 82.0642
500 ng/l Spike B | Chlor | 5.8409 | 1091.3199
.2 m nylon + 1/31filt + 50ng/l B | Chlor | 1.5213 | 21.8269
.2 m nylon + 1/31filt + 50ng/l B | Chlor | 1.3528 | 18.2607
.2 m nylon + Disposable syr. + 50ng/l B | Chlor | 0.28824 | 12.6455
.2 m nylon + Disposable syr. + 50ng/l B | Chlor | 0.39978 | 9.7994
.45 m nylon + 1/31filt + 50ng/l B | Chlor | 12.5326 | 34.3261
.45 m nylon + 1/31filt + 50ng/l B | Chlor | 2.1035 | 26.3692
1m glass filt. + 1/31 filt. + 50ng/l B | Chlor | 1.4587 | 41.5905
1m glass filt. + 1/31 filt. + 50ng/l B | Chlor | 1.4517 | 40.7918
.2 m nylon + 1/31filt + 500ng/l B | Chlor | 2.1681 | 146.5363
.2 m nylon + 1/31filt + 500ng/l B | Chlor | 2.1153 | 155.6147
.2 m nylon + Disposable syr. + 500ng/l B | Chlor | 5.4515 | 13.3123
.2 m nylon + Disposable syr. + 500ng/l B | Chlor | 7.7477 | 11.7131
.45 m nylon + 1/31filt + 500ng/l B | Chlor | 1.6948 | 266.5243
.45 m nylon + 1/31filt + 500ng/l B | Chlor | 16.7664 | 247.8321
1m glass filt. + 1/31 filt. + 500ng/l B | Chlor | 5.8353 | 479.1545
1m glass filt. + 1/31 filt. + 500ng/l B | Chlor | 7.1616 | 414.704
MDL 40ng/l B | Chlor | 1.0193 | 53.6164
MDL 40ng/l B | Chlor | 2.3174 | 42.6177
MDL 40ng/l B | Chlor | 2.0053 | 47.643
MDL 40ng/l B | Chlor | 1.3172 | 37.7125
MDL 40ng/l B | Chlor | 14.1847 | 52.9076
QA SPIKE (nominal spike by ME) C | chlor | 3.7736 | 137.4813
MDL 40 (nominal 40 ppt spike) C | chlor | 4.3995 | 14.0941
QA evaluation sample C | chlor | 3.5293 | 201.9737
P & A 50 ppt C | chlor | 7.4272 | 52.5908
P & A 125 ppt C | chlor | 0.65603 | 177.4843
P & A 300 ppt C | chlor | 2.9707 | 366.63 | Sample Description | Nominal | Measured | % RECOVERY
P & A 50 ppt C | chlor | 0.77449 | 55.8718 |  | Concentration | Concentration
P & A 125 ppt C | chlor | 6.7387 | 168.8185 |  | (μg/L) | (μg/L)
P & A 300 ppt C | chlor | 16.9809 | 331.4542 | .050 μg/L Calibrator | 0.05 | 0.057315 | 114.6293
MDL 40ng/l C | chlor | 15.6097 | 71.2276 | .125 μg/L Calibrator | 0.125 | 0.14978 | 119.8259
MDL 40ng/l C | chlor | 13.4175 | 36.5502 | .300 μg/L Calibrator | 0.25 | 0.19728 | 78.912
MDL 40ng/l C | chlor | 7.9157 | 34.7916 | .050 μg/L Calibrator | 0.03 | 0.052591 | 175.3028
MDL 40ng/l C | chlor | 0.56883 | 34.2244 | .125 μg/L Calibrator | 0.1 | 0.17748 | 177.4843
MDL 40ng/l C | chlor | 11.3483 | 43.1766 | .300 μg/L Calibrator | 0.25 | 0.36663 | 146.652
MDL 40ng/l C | chlor | 28.8578 | 29.3861 | .050 μg/L Calibrator | 0.05 | 0.055872 | 111.7437
50 ng/l Spike C | chlor | 9.3044 | 2.1406 | .125 μg/L Calibrator | 0.125 | 0.16882 | 135.0548
500 ng/l Spike C | chlor | 1.6248 | 8.561 | .300 μg/L Calibrator | 0.3 | 0.33145 | 110.4847
.2 m nylon + 1/31filt + 50ng/l C | chlor | 0.39995 | 22.012
.2 m nylon + Disposable syr. + 50ng/l C | chlor | 0.10138 | 9.4958
.45 m nylon + 1/31filt + 50ng/l C | chlor | 2.8338 | 9.4101
1m glass filt. + 1/31 filt. + 50ng/l C | chlor | 5.2946 | 11.165
.2 m nylon + 1/31filt + 500ng/l C | chlor | 11.2061 | 4.3762
.2 m nylon + Disposable syr. + 500ng/l C | chlor | 15.3473 | 7.4048
.45 m nylon + 1/31filt + 500ng/l C | chlor | 3.131 | 16.7042
1m glass filt. + 1/31 filt. + 500ng/l C | chlor | 0.64395 | 53.1849
50 ng/l Spike DUPLICATE C | chlor | 3.795 | 9.1813
500 ng/l Spike DUPLICATE C | chlor | 7.1503 | 4.7055
.2 m nylon + 1/31filt + 50ng/l C | chlor | 19.7947 | 5.1923
50 ppt@100% C | chlor | 0.91904 | 60.094
500ppt@50% C | chlor | 7.0157 | 223.0866
500ppt@25% C | chlor | 4.3463 | 88.9403
50ppt C | chlor | 4.5402 | 10.9654
500ppt C | chlor | 2.9662 | 10.4576
50ppt Dup. C | chlor | 2.5682 | 90.8432
500ppt Dup. C | chlor | 3.4685 | 341.8367
50 ppt@100% Dup. C | chlor | 0.70593 | 203.9352
500ppt@50% Dup. C | chlor | 11.7948 | 25.8337
500ppt@25% Dup. C | chlor | 7.1049 | 45.3285
QA FOR ANALYTE: DIAZINON
Standard 30 A | diaz | 2.4456 | 28.096
Standard 100 A | diaz | 2.5986 | 92.1102
Standard 250 A | diaz | 0.25791 | 264.4977
Method Blank #2 A | diaz | 1.0099 | 16.0451
Method Blank #3 A | diaz | 4.0699 | 14.5224
Method Blank #4 A | diaz | 0.81319 | 15.1252
Evaluation Sample A | diaz | 0.25269 | 252.9582
50 ng/l Spike B | diaz | 5.2378 | 200.525
500 ng/l Spike B | diaz | 5.3289 | 636.4161
50 ppt filter QA spike C | diaz | 0.98396 | 39.2795
500 ppt filter QA spike @50% C | diaz | 4.8128 | 45.128
.2 m nylon + 1/31filt + 50ng/l B | diaz | 3.4493 | 44.0191
.2 m nylon + 1/31filt + 50ng/l B | diaz | 1.3622 | 62.8523
.2 m nylon + Disposable syr. + 50ng/l B | diaz | 1.4535 | 79.1422
.2 m nylon + Disposable syr. + 50ng/l B | diaz | 8.4088 | 105.8411
.45 m nylon + 1/31filt + 50ng/l B | diaz | 1.0904 | 49.6546
.45 m nylon + 1/31filt + 50ng/l B | diaz | 2.376 | 50.1118
1m glass filt. + 1/31 filt. + 50ng/l B | diaz | 4.2788 | 456.3493
1m glass filt. + 1/31 filt. + 50ng/l B | diaz | 0.10971 | 456.3493
.2 m nylon + 1/31filt + 500ng/l B | diaz | 0.7801 | 468.4583
.2 m nylon + 1/31filt + 500ng/l B | diaz | 7.1219 | 575.3778
.2 m nylon + Disposable syr. + 500ng/l B | diaz | 6.7865 | 454.5601
.2 m nylon + Disposable syr. + 500ng/l B | diaz | 2.4148 | 422.9744
.45 m nylon + 1/31filt + 500ng/l B | diaz | 1.0158 | 478.3764
.45 m nylon + 1/31filt + 500ng/l B | diaz | 2.0629 | 490.4272
1m glass filt. + 1/31 filt. + 500ng/l B | diaz | 0.76314 | 21.109
1m glass filt. + 1/31 filt. + 500ng/l B | diaz | 1.689 | 23.8889
MDL 20ng/l B | diaz | 2.1978 | 20.5816
MDL 20ng/l B | diaz | 0.31906 | 20.38
MDL 20ng/l B | diaz | 2.3884 | 25.9183
50 ppt filter QA spike C | diaz | 2.0289 | 29.4354
500 ppt filter QA spike @50% C | diaz | 0.63992 | 267.9785

## Diaz.PA
QA FOR ANALYTE: DIAZINON | %CV | ng/L | % Recovery
P&A Standard 30 A | 2.4456 | 28.096 | 93.6667 |  | Sample Description | Nominal | Measured | % RECOVERY
P&A Standard 100 A | 2.5986 | 92.1102 | 92.11 |  |  | Concentration | Concentration
P&A Standard 250 A | 0.25791 | 264.4977 | 105.8 |  |  | (μg/L) | (μg/L)
[5 empty] | 0.030 μg/L Calibrator | 0.03 | 0.028 | 93.3333
[5 empty] | 0.100 μg/L Calibrator | 0.1 | 0.092 | 92
|  |  |  | Yellow is below linear range of 29.53 | 0.250 μg/L Calibrator | 0.25 | 0.264 | 105.6
Key
02/21/2001=A
02/23/2001=B
02/28/2001=C

## Diaz.Method Blank
QA FOR ANALYTE: DIAZINON | %CV | ng/L
Method Blank #2 A | 1.0099 | 16.0451
Method Blank #3 A | 4.0699 | 14.5224
Method Blank #4 A | 0.81319 | 15.1252
Yellow is below linear range of 29.53
Evaluation Sample A | 0.25269 | 252.9582
Key
02/21/2001=A
02/23/2001=B
02/28/2001=C

## Diaz.Filters
QA FOR ANALYTE: DIAZINON | %CV | ng/L
50 ng/l Spike B | 5.2378 | 200.525
500 ng/l Spike B | 5.3289 | 636.4161
50 ppt filter QA spike C | 2.0289 | 29.4354
500 ppt filter QA spike @50% C | 0.63992 | 267.9785
.2 m nylon + 1/31filt + 50ng/l B | 0.98396 | 39.2795
.2 m nylon + 1/31filt + 50ng/l B | 4.8128 | 45.128
.2 m nylon + Disposable syr. + 50ng/l B | 3.4493 | 44.0191
.2 m nylon + Disposable syr. + 50ng/l B | 1.3622 | 62.8523
.45 m nylon + 1/31filt + 50ng/l B | 1.4535 | 79.1422
.45 m nylon + 1/31filt + 50ng/l B | 8.4088 | 105.8411
1m glass filt. + 1/31 filt. + 50ng/l B | 1.0904 | 49.6546
1m glass filt. + 1/31 filt. + 50ng/l B | 2.376 | 50.1118
.2 m nylon + 1/31filt + 500ng/l B | 4.2788 | 456.3493
.2 m nylon + 1/31filt + 500ng/l B | 0.10971 | 456.3493
.2 m nylon + Disposable syr. + 500ng/l B | 0.7801 | 468.4583
.2 m nylon + Disposable syr. + 500ng/l B | 7.1219 | 575.3778
.45 m nylon + 1/31filt + 500ng/l B | 6.7865 | 454.5601
.45 m nylon + 1/31filt + 500ng/l B | 2.4148 | 422.9744
1m glass filt. + 1/31 filt. + 500ng/l B | 1.0158 | 478.3764
1m glass filt. + 1/31 filt. + 500ng/l B | 2.0629 | 490.4272
Yellow is out of linear range of 458.15
Key
02/21/2001=A
02/23/2001=B
02/28/2001=C

## Diaz.MDL
QA FOR ANALYTE: DIAZINON | %CV | ng/L | % Recovery | actual Conc. | Found Concentration | % Recovery
MDL 20ng/l B | 0.76314 | 21.109 | 105.55 | .02 ug/l | 21.109 | 105
MDL 20ng/l B | 1.689 | 23.8889 | 119.45 | .02 ug/l | 23.8889 | 120
MDL 20ng/l B | 2.1978 | 20.5816 | 102.9 | .02 ug/l | 20.5816 | 105
MDL 20ng/l B | 0.31906 | 20.38 | 101.9 | .02 ug/l | 20.38 | 100
MDL 20ng/l B | 2.3884 | 25.9183 | 129.6 | .02 ug/l | 25.9183 | 130
|  |  |  | Average |  | 22.3756
|  |  |  | STDEV |  | 2.4314
|  |  |  | MDL=2.896 x STDEV |  | 5.792
Key
02/21/2001=A
02/23/2001=B
02/28/2001=C |  | Yellow is below linear range of 29.09

## Chlor.PA
QA FOR ANALYTE: CHLORPYRIFOS | %CV | ng/L | % Recovery | Minus High CV
Standard 50 A | 4.9781 | 57.3146 | 114.62 | 114.62
Standard 125 A | 10.4837 | 149.7823 | 119.824
Standard 300 A | 4.5051 | 197.2799 | 65.76 | 65.76
50 ppt C | 7.4272 | 52.5908 | 105.18 | 105.18
125 ppt C | 0.65603 | 177.4843 | 141.984 | 141.984
300 ppt C | 2.9707 | 366.63 | 122.21 | 122.21
50 ppt C | 0.77449 | 55.8718 | 111.74 | 111.74
125 ppt C | 6.7387 | 168.8185 | 135.056 | 135.056
300 ppt C | 16.9809 | 331.4542 | 110.4833
Average | 6.1683 | 173.0252 | 114.0953 | 113.7929
STDEV | 5.1397 | 114.7644 | 21.6519 | 24.846
Key
02/21/2001=A
02/23/2001=B
02/28/2001=C

## Chlor.Method Blank
QA FOR ANALYTE: CHLORPYRIFOS | %CV | ng/L
Method Blank #2 A | 3.2841 | 27.8321
Method Blank #3 A | 1.5268 | 27.9167
Method Blank #4 A | 2.5963 | 26.0541
Method Blank #1 A | 2.0049 | 21.915
Yellow is below linear range of 49.81
Key
02/21/2001=A
02/23/2001=B
02/28/2001=C

## Chlor.Filters
QA FOR ANALYTE: CHLORPYRIFOS | %CV | ng/L
50 ng/l Spike B | 5.6275 | 82.0642
500 ng/l Spike B | 5.8409 | 1091.3199 | Key
.2 m nylon + 1/31filt + 50ng/l B | 1.5213 | 21.8269 | 02/21/2001=A
.2 m nylon + 1/31filt + 50ng/l B | 1.3528 | 18.2607 | 02/23/2001=B
.2 m nylon + Disposable syr. + 50ng/l B | 0.28824 | 12.6455 | 02/28/2001=C
.2 m nylon + Disposable syr. + 50ng/l B | 0.39978 | 9.7994
.45 m nylon + 1/31filt + 50ng/l B | 12.5326 | 34.3261
.45 m nylon + 1/31filt + 50ng/l B | 2.1035 | 26.3692
1m glass filt. + 1/31 filt. + 50ng/l B | 1.4587 | 41.5905
1m glass filt. + 1/31 filt. + 50ng/l B | 1.4517 | 40.7918
.2 m nylon + 1/31filt + 500ng/l B | 2.1681 | 146.5363
.2 m nylon + 1/31filt + 500ng/l B | 2.1153 | 155.6147
.2 m nylon + Disposable syr. + 500ng/l B | 5.4515 | 13.3123
.2 m nylon + Disposable syr. + 500ng/l B | 7.7477 | 11.7131
.45 m nylon + 1/31filt + 500ng/l B | 1.6948 | 266.5243
.45 m nylon + 1/31filt + 500ng/l B | 16.7664 | 247.8321
1m glass filt. + 1/31 filt. + 500ng/l B | 5.8353 | 479.1545
1m glass filt. + 1/31 filt. + 500ng/l B | 7.1616 | 414.704
50 ng/l Spike C | 9.3044 | 2.1406
500 ng/l Spike C | 1.6248 | 8.561
50 ng/l Spike Duplicate C | 3.795 | 9.1813
500 ng/l Spike Duplicate C | 7.1503 | 4.7055
.2 m nylon + 1/31filt + 50ng/l C | 0.39995 | 22.012
.2 m nylon + Disposable syr. + 50ng/l C | 0.10138 | 9.4958
.45 m nylon + 1/31filt + 50ng/l C | 2.8338 | 9.4101
1m glass filt. + 1/31 filt. + 50ng/l C | 5.2946 | 11.165
.2 m nylon + 1/31filt + 500ng/l C | 11.2061 | 4.3762
.2 m nylon + Disposable syr. + 500ng/l C | 15.3473 | 7.4048
.45 m nylon + 1/31filt + 500ng/l C | 3.131 | 16.7042
1m glass filt. + 1/31 filt. + 500ng/l C | 0.64395 | 53.1849
.2 m nylon + 1/31filt + 50ng/l C | 19.7947 | 5.1923
50 ppt@100% 2/23 B | 0.91904 | 60.094
500ppt@50% 2/23 B | 7.0157 | 223.0866
500ppt@25% 2/23B | 4.3463 | 88.9403
50ppt 2/28 C | 4.5402 | 10.9654
500ppt 2/28 C | 2.9662 | 10.4576
50ppt 2/28 Dup. C | 11.7948 | 25.8337
500ppt 2/28 Dup. C | 7.1049 | 45.3285
50 ppt@100% 2/23 Dup. B | 2.5682 | 90.8432
500ppt@50% 2/23 Dup. B | 3.4685 | 341.8367
500ppt@25% 2/23 Dup. B | 0.70593 | 203.9352

## Chlor.MDL
QA FOR ANALYTE: CHLORPYRIFOS | %CV | ng/L | actual Conc. | Found Concentration | % Recovery | Minus High CV
MDL 40ng/l B | 1.0193 | 53.6164 | 0.04ug/L | 53.6164 | 134.05 | 53.6164
MDL 40ng/l B | 2.3174 | 42.6177 | 0.04ug/L | 42.6177 | 106.55 | 42.6177
MDL 40ng/l B | 2.0053 | 47.643 | 0.04ug/L | 47.643 | 119.1 | 47.643
MDL 40ng/l B | 1.3172 | 37.7125 | 0.04ug/L | 37.7125 | 94.275 | 37.7125
MDL 40ng/l B | 14.1847 | 52.9076 | 0.04ug/L | 52.9076 | 132.275
MDL 40ng/l C | 15.6097 | 71.2276 | 0.04ug/L | 71.2276 | 178.075
MDL 40ng/l C | 13.4175 | 36.5502 | 0.04ug/L | 36.5502 | 91.375
MDL 40ng/l C | 7.9157 | 34.7916 | 0.04ug/L | 34.7916 | 86.975 | 34.7916
MDL 40ng/l C | 0.56883 | 34.2244 | 0.04ug/L | 34.2244 | 85.55 | 34.2244
MDL 40ng/l C | 11.3483 | 43.1766 | 0.04ug/L | 43.1766 | 107.95
MDL 40ng/l C | 28.8578 | 29.3861 | 0.04ug/L | 29.3861 | 73.475
|  |  | Average | 41.5013 | 109.9682 | 53.6164
|  |  | STDEV | 11.9079 | 28.3867 | 7.7155
|  |  | MDL=2.896 x STDEV | 34.4852 | 82.208 | 22.3442
1SPIKE (nominal spike by ME) C | 3.7736 | 137.4813
2MDL 40 (nominal 40 ppt spike) C | 4.3995 | 14.0941
Evaluation sample C | 3.5293 | 201.9737
Yellow is above linear range of 43.79
1Lab QA officer spiked contol water with a measured concentration = 137
2Minimum detection limit nominal spike = 40ng/L